| Photometric Test Report |  |  |
| --- | --- | --- |
| Report Number: POTS/GJ13036 | Report Date: 01-02-2013 | Prepared By: G John |
| Test Laboratory: Photometric and Optical Testing Services, Cotswold Business Centre, 42 A P Ellis Road, Rissington Business Park, Upper Rissington, Gloucestershire, GL54 2QB |  |  |
| Company Registration Number: Registered in England & Wales No. OC352911 |  |  |
| Registered Address: Thistle Down Barn, Holcot Lane, Sywell, Northampton, NN6 0BG |  |  |

| Client Details |  |
| --- | --- |
| Company: Astro Lighting | Email: technical@astrolighting.co.uk |
| Address: Astro Lighting Limited, G2 River Way, Harlow CM20 2DP, Great Britain |  |

| Details of Product Tested |  |
| --- | --- |
| Manufacturer: Astro Lighting | Source Type: Domestic light |
| Model: FUSE UNSWITCHED BRONZE | Serial Number: 7046 |
| Lamp Type: LED |  |
| Power Supply Used: Uninterruptible AC power supply |  |
| Voltage(AC V) = 246.3 | Current (mA)= 31 |
| Power (Watts)= 3.86 | Power factor= 0.505 |

| Integrating Sphere Test |  |  |
| --- | --- | --- |
| Date of Test: 08-01-2013 |  | Ambient Temperature:25°C |
| Measurement Filename: 7046 FUSE UNSWITCHED BRONZE |  |  |
| Instrument Used: Labsphere model CSLMS LED 4060 integrating sphere spectroradiometer |  |  |
| Integrating Sphere Size: 1m |  | Measurement Geometry (2 / 4): 4 |
| Sample Orientation: Horizontal |  | Auxiliary Correction Applied: YES |
| Comments: |  |  |
| Date of Last Calibration (Operating Hours): 08-11-2012 (02:36) |  | Spectral Flux Standard Lamp Used: SCL-1400 |
| Standard Lamp Serial Number: K75 |  | Traceable: to NIST standards |
| Calibration Certificate Number: DM-02008-001 |  | Calibration Certificate Date: 19th February 2010 |
| Calibration Lamp Uncertainty: ± 0.67% (k=2) |  |  |
| Results |  |  |
| Flux (lumens): 80.88 |  |  |
| CIE 1931 Chromaticity Cx: 0.4393 |  | CIE 1931 Chromaticity Cy: 0.4079 |
| CRI (%):82.4 |  | CCT (K): 2990 |

Figure 1: Spectral Irradiance

Figure 2: CIE 1931 diagram.

| Goniophotometer Test |  |  |  |
| --- | --- | --- | --- |
| Date of Test: 10-01-2013 |  | Ambient Temperature: 25°C |  |
| Measurement Filename: 7046 FUSE UNSWITCHED BRONZE |  |  |  |
| Instrument Used: Radiant Imaging NFMS0800 Goniometer with ProMetric PM-1423F-1 Imaging Photometer |  |  |  |
| Photometer Working Distance: 2 m |  | Measurement Geometry: Near-Field |  |
| Comments: |  |  |  |
| Reference Photometer Used: Specbos1201 |  | Reference Photometer Serial Number: 2911670 |  |
| Traceable: to NPL standards, UKAS Accredited |  | Calibration Certificate Number: 121104 |  |
| Calibration Certificate Date: 25th October 2011 |  | Sample Stabilisation Time (minutes):25 |  |
| Reference Photometer Calibration Uncertainty: ± 2.4% (k=2, 20-200 lux, CIE illuminant A source) |  |  |  |
| Scan Set Up |  |  |  |
| Direction | Range |  | Increment |
| Inclination Zone 1 | 0-90° |  | 3° |
| Azimuth | 0-360° |  | 10° |
| Results |  |  |  |
| Integrated Luminous Flux (lumens):80.88 | Peak Intensity (3° Spot, candelas): 411.1 |  | Efficacy (lumens/Watt): 20.95 |
| Beam Angle (50% of max intensity C0-180, degrees): 16.0 |  |  |  |
| Photometric Filename (IES LM-63-2002): 7046 FUSE UNSWITCHED BRONZE |  |  |  |
| IES File – Absolute or Relative Format? Absolute |  |  |  |
| Photometric Filename (EULUMDAT): 7046 FUSE UNSWITCHED BRONZE |  |  |  |
| EULUMDAT File – Absolute or Relative Format? Absolute |  |  |  |

Figure 3: Far-Field Luminous Intensity (C0-180, Cartesian Coordinates)

Figure 4: Far-Field Luminous Intensity (C0-180, C90-270, Polar Coordinates)

Figure 4. Illuminance cone diagram.

Table 1. UGR values

|  | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 |
| 3 | 402.1 | 409.2 | 411.1 | 406.6 | 408.1 | 407.4 | 399.4 | 394.7 | 386.6 | 378.5 | 367.4 | 364.8 | 347.8 | 348.3 | 343.5 | 333.1 | 334.5 | 335.7 | 329.8 |
| 6 | 305.7 | 299.3 | 300.4 | 308.0 | 289.1 | 293.0 | 293.8 | 271.7 | 268.1 | 270.1 | 249.1 | 238.0 | 240.2 | 223.9 | 216.8 | 227.2 | 214.6 | 214.5 | 220.3 |
| 9 | 199.5 | 209.7 | 195.1 | 203.4 | 198.7 | 191.6 | 199.6 | 178.5 | 185.7 | 174.8 | 170.4 | 172.3 | 157.3 | 169.8 | 163.1 | 157.7 | 162.3 | 148.6 | 154.0 |
| 12 | 136.2 | 137.6 | 130.4 | 143.2 | 131.5 | 142.0 | 135.0 | 133.4 | 124.0 | 122.7 | 118.6 | 118.3 | 123.3 | 114.5 | 122.3 | 113.1 | 116.1 | 103.4 | 107.2 |
| 15 | 91.8 | 95.3 | 90.8 | 100.6 | 94.6 | 98.3 | 96.7 | 89.3 | 86.9 | 84.8 | 86.4 | 82.6 | 87.0 | 81.5 | 83.9 | 80.1 | 78.2 | 72.0 | 75.7 |
| 18 | 62.0 | 63.6 | 62.2 | 61.0 | 65.9 | 63.1 | 60.9 | 58.8 | 58.7 | 58.7 | 55.2 | 59.3 | 52.3 | 55.8 | 54.4 | 51.3 | 51.5 | 49.0 | 51.9 |
| 21 | 41.4 | 40.6 | 39.4 | 39.5 | 40.4 | 42.1 | 39.7 | 39.7 | 40.2 | 39.2 | 37.9 | 38.0 | 34.6 | 35.2 | 34.2 | 35.1 | 33.7 | 33.2 | 35.2 |
| 24 | 29.0 | 26.1 | 26.0 | 27.6 | 28.1 | 29.7 | 28.4 | 28.9 | 27.8 | 27.3 | 26.6 | 26.5 | 25.3 | 25.8 | 23.8 | 24.6 | 23.3 | 22.7 | 25.1 |
| 27 | 19.0 | 19.7 | 19.9 | 19.4 | 22.1 | 20.2 | 20.5 | 21.7 | 19.8 | 20.0 | 19.8 | 19.2 | 19.3 | 18.8 | 17.3 | 18.6 | 17.4 | 17.0 | 18.3 |
| 30 | 14.1 | 14.9 | 15.3 | 15.1 | 15.0 | 15.8 | 14.9 | 14.4 | 15.5 | 14.7 | 14.8 | 15.0 | 15.1 | 13.6 | 13.4 | 13.8 | 12.4 | 13.1 | 13.9 |
| 33 | 11.7 | 11.0 | 10.4 | 11.1 | 12.0 | 10.9 | 10.7 | 11.6 | 11.1 | 10.6 | 11.2 | 11.8 | 10.4 | 10.5 | 10.3 | 10.2 | 9.8 | 10.5 | 10.9 |
| 36 | 7.7 | 7.7 | 8.9 | 8.9 | 8.2 | 8.5 | 9.0 | 8.7 | 8.2 | 8.6 | 8.9 | 8.9 | 8.2 | 8.0 | 8.1 | 7.9 | 7.6 | 8.0 | 8.2 |
| 39 | 6.9 | 6.8 | 6.4 | 6.1 | 7.0 | 7.0 | 6.8 | 6.4 | 6.7 | 7.1 | 7.2 | 6.9 | 6.9 | 6.5 | 6.4 | 6.2 | 6.1 | 6.5 | 6.9 |
| 42 | 5.0 | 5.3 | 5.3 | 5.0 | 5.0 | 5.3 | 5.3 | 5.3 | 5.6 | 5.5 | 5.7 | 5.4 | 5.8 | 5.7 | 5.5 | 5.1 | 5.5 | 5.6 | 5.6 |
| 45 | 4.4 | 4.2 | 4.4 | 4.2 | 4.1 | 4.5 | 4.6 | 4.3 | 4.5 | 4.6 | 4.7 | 4.9 | 4.9 | 4.6 | 4.4 | 4.4 | 4.5 | 4.9 | 5.1 |
| 48 | 3.6 | 3.6 | 3.3 | 3.6 | 3.9 | 3.9 | 3.6 | 3.9 | 3.9 | 3.7 | 3.9 | 4.0 | 4.5 | 4.3 | 4.1 | 4.1 | 4.5 | 4.5 | 4.7 |
| 51 | 3.2 | 3.2 | 3.3 | 3.2 | 3.5 | 3.3 | 3.4 | 3.6 | 3.5 | 3.4 | 3.7 | 3.7 | 4.2 | 3.9 | 4.0 | 3.7 | 4.0 | 4.3 | 4.3 |
| 54 | 2.8 | 3.0 | 3.0 | 2.9 | 3.0 | 3.2 | 3.4 | 3.5 | 3.4 | 3.6 | 4.0 | 4.0 | 3.5 | 3.9 | 3.8 | 3.7 | 3.8 | 4.3 | 4.3 |
| 57 | 3.0 | 3.2 | 3.1 | 3.0 | 2.9 | 2.9 | 3.2 | 3.2 | 3.2 | 3.6 | 3.7 | 3.9 | 3.8 | 3.9 | 3.7 | 3.7 | 3.9 | 4.3 | 4.1 |
| 60 | 3.0 | 3.0 | 3.2 | 2.8 | 3.0 | 3.2 | 3.3 | 3.1 | 3.3 | 3.0 | 3.6 | 3.8 | 3.8 | 3.8 | 3.9 | 3.9 | 3.7 | 4.0 | 4.1 |
| 63 | 2.9 | 3.2 | 3.0 | 2.9 | 2.9 | 3.2 | 3.1 | 3.1 | 3.1 | 3.3 | 3.2 | 3.8 | 3.8 | 3.6 | 3.8 | 4.0 | 3.9 | 3.7 | 4.0 |
| 66 | 2.8 | 3.0 | 2.7 | 2.8 | 2.9 | 3.1 | 3.1 | 2.9 | 3.3 | 3.0 | 3.2 | 3.6 | 3.7 | 3.6 | 3.6 | 3.7 | 3.6 | 3.8 | 3.7 |
| 69 | 2.9 | 2.9 | 2.9 | 3.2 | 2.8 | 3.0 | 2.9 | 3.3 | 3.1 | 3.1 | 3.3 | 3.5 | 3.5 | 3.6 | 3.5 | 3.5 | 3.4 | 3.6 | 3.8 |
| 72 | 2.7 | 2.6 | 2.7 | 2.6 | 2.6 | 2.9 | 3.0 | 3.1 | 3.0 | 3.0 | 3.3 | 3.6 | 3.7 | 3.6 | 3.7 | 3.8 | 3.7 | 3.7 | 3.8 |
| 75 | 2.8 | 2.5 | 2.7 | 2.8 | 2.5 | 2.8 | 2.8 | 3.2 | 3.1 | 3.1 | 3.3 | 3.5 | 3.4 | 3.7 | 3.5 | 3.5 | 3.5 | 3.7 | 3.9 |
| 78 | 3.1 | 2.6 | 2.7 | 2.7 | 2.6 | 2.8 | 3.1 | 2.9 | 3.1 | 3.2 | 3.4 | 3.1 | 3.6 | 3.7 | 3.6 | 3.6 | 3.6 | 3.5 | 4.0 |
| 81 | 3.1 | 2.6 | 2.6 | 2.8 | 2.9 | 2.7 | 2.8 | 3.0 | 2.8 | 3.0 | 3.1 | 3.1 | 3.5 | 3.6 | 3.6 | 3.4 | 3.5 | 3.7 | 4.3 |
| 84 | 3.2 | 2.7 | 2.6 | 2.7 | 2.8 | 2.7 | 2.7 | 3.0 | 3.4 | 3.1 | 3.3 | 3.4 | 3.7 | 3.5 | 3.6 | 3.3 | 3.6 | 3.8 | 4.3 |
| 87 | 2.6 | 2.1 | 2.5 | 2.4 | 2.5 | 2.6 | 2.8 | 2.8 | 2.7 | 2.7 | 3.2 | 3.2 | 3.0 | 3.4 | 3.3 | 3.2 | 3.1 | 3.0 | 3.9 |
| 90 | 0.9 | 0.9 | 0.6 | 0.9 | 1.3 | 0.9 | 0.9 | 0.6 | 1.3 | 1.1 | 1.5 | 0.8 | 1.0 | 1.0 | 1.5 | 1.2 | 0.9 | 1.6 | 1.6 |

Table 2a. Luminous intensity values, azimuth 0-180°

|  | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | 310 | 320 | 330 | 340 | 350 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 | 408.6 |
| 3 | 335.1 | 330.6 | 326.2 | 334.6 | 339.4 | 338.8 | 347.5 | 349.4 | 356.0 | 363.0 | 368.2 | 373.0 | 378.5 | 386.6 | 393.0 | 397.0 | 399.3 |
| 6 | 220.2 | 210.6 | 226.0 | 226.0 | 225.9 | 241.4 | 239.1 | 241.6 | 253.6 | 251.9 | 258.1 | 273.4 | 270.6 | 284.4 | 299.3 | 289.5 | 304.6 |
| 9 | 146.8 | 151.0 | 154.4 | 153.0 | 163.2 | 158.2 | 163.7 | 165.4 | 166.2 | 168.8 | 164.3 | 182.4 | 181.5 | 195.8 | 198.3 | 199.5 | 207.7 |
| 12 | 98.8 | 106.0 | 105.3 | 108.2 | 110.4 | 113.4 | 116.1 | 114.7 | 116.4 | 116.0 | 116.5 | 122.8 | 126.4 | 127.6 | 138.8 | 130.7 | 138.2 |
| 15 | 71.1 | 76.2 | 75.1 | 77.6 | 76.5 | 83.3 | 80.5 | 82.6 | 83.8 | 84.6 | 86.1 | 89.9 | 91.4 | 91.6 | 98.7 | 89.9 | 95.9 |
| 18 | 48.3 | 53.0 | 49.5 | 50.4 | 52.8 | 53.0 | 56.9 | 56.9 | 59.3 | 61.6 | 63.4 | 61.1 | 62.3 | 63.3 | 64.7 | 64.1 | 65.0 |
| 21 | 32.1 | 34.8 | 33.2 | 33.1 | 34.3 | 33.3 | 37.9 | 37.5 | 41.0 | 41.6 | 40.5 | 43.1 | 41.0 | 41.7 | 43.2 | 42.3 | 44.5 |
| 24 | 22.5 | 24.0 | 24.3 | 24.1 | 24.7 | 23.8 | 25.5 | 26.6 | 26.8 | 28.1 | 27.8 | 28.2 | 30.9 | 27.5 | 30.0 | 28.4 | 28.9 |
| 27 | 16.5 | 18.0 | 18.1 | 17.0 | 18.2 | 18.6 | 17.8 | 20.3 | 19.4 | 18.9 | 21.9 | 19.6 | 19.5 | 20.8 | 19.9 | 20.9 | 20.5 |
| 30 | 12.7 | 12.7 | 14.3 | 13.5 | 13.1 | 14.9 | 14.2 | 13.3 | 14.7 | 15.3 | 13.9 | 14.1 | 15.3 | 14.5 | 14.7 | 15.4 | 15.3 |
| 33 | 10.2 | 9.7 | 10.5 | 10.9 | 10.5 | 10.8 | 11.6 | 10.9 | 10.4 | 11.4 | 12.0 | 10.5 | 10.5 | 11.6 | 11.2 | 10.2 | 10.7 |
| 36 | 8.2 | 7.6 | 8.1 | 8.7 | 8.2 | 8.0 | 8.6 | 8.9 | 8.4 | 8.3 | 8.7 | 8.8 | 8.1 | 7.8 | 8.4 | 9.2 | 8.4 |
| 39 | 6.6 | 6.2 | 6.6 | 6.9 | 6.9 | 6.8 | 6.5 | 6.9 | 7.3 | 6.7 | 6.5 | 6.5 | 6.6 | 6.7 | 6.3 | 6.2 | 6.4 |
| 42 | 5.5 | 5.5 | 5.4 | 6.1 | 5.9 | 5.7 | 5.6 | 5.3 | 5.5 | 5.8 | 5.4 | 5.2 | 5.2 | 5.0 | 5.3 | 5.4 | 5.2 |
| 45 | 4.7 | 4.6 | 4.8 | 5.4 | 5.0 | 4.8 | 4.7 | 4.6 | 4.3 | 4.6 | 4.4 | 4.6 | 4.6 | 4.2 | 4.4 | 4.3 | 4.3 |
| 48 | 4.3 | 4.3 | 4.4 | 4.4 | 4.6 | 4.2 | 4.5 | 4.2 | 4.3 | 4.0 | 4.0 | 3.6 | 3.7 | 4.0 | 3.8 | 3.6 | 3.5 |
| 51 | 4.3 | 4.5 | 4.4 | 4.3 | 4.2 | 4.3 | 4.0 | 3.9 | 4.2 | 4.0 | 3.7 | 3.4 | 3.4 | 3.5 | 3.4 | 3.3 | 3.2 |
| 54 | 4.0 | 4.2 | 4.1 | 4.2 | 4.1 | 4.0 | 4.0 | 4.0 | 3.8 | 3.8 | 3.3 | 3.5 | 3.4 | 3.3 | 2.8 | 3.2 | 3.1 |
| 57 | 4.2 | 3.9 | 4.1 | 3.9 | 4.1 | 4.1 | 4.1 | 3.7 | 3.8 | 3.6 | 3.2 | 3.2 | 3.3 | 3.0 | 3.1 | 3.2 | 3.2 |
| 60 | 3.9 | 3.7 | 3.8 | 4.0 | 4.0 | 4.1 | 3.9 | 3.8 | 3.7 | 3.7 | 3.3 | 3.1 | 3.1 | 3.1 | 3.2 | 3.0 | 2.9 |
| 63 | 3.8 | 3.9 | 3.4 | 3.9 | 3.9 | 3.7 | 3.6 | 3.5 | 3.4 | 3.3 | 3.2 | 3.1 | 3.2 | 3.2 | 3.1 | 3.0 | 3.1 |
| 66 | 3.8 | 3.5 | 3.6 | 3.6 | 4.0 | 3.7 | 3.7 | 3.4 | 3.3 | 3.3 | 3.1 | 3.2 | 2.9 | 3.0 | 2.9 | 2.9 | 2.7 |
| 69 | 3.7 | 3.9 | 3.6 | 3.6 | 3.9 | 3.6 | 3.2 | 3.4 | 3.3 | 3.1 | 2.9 | 2.9 | 2.9 | 3.0 | 3.0 | 2.7 | 2.9 |
| 72 | 3.8 | 3.6 | 3.6 | 3.7 | 3.7 | 3.5 | 3.1 | 3.6 | 3.2 | 3.3 | 3.0 | 2.6 | 2.8 | 2.7 | 3.0 | 3.0 | 2.7 |
| 75 | 3.5 | 3.3 | 3.8 | 3.4 | 3.5 | 3.2 | 3.3 | 3.3 | 3.4 | 3.4 | 2.8 | 3.0 | 3.1 | 2.8 | 3.1 | 2.7 | 2.8 |
| 78 | 3.5 | 3.7 | 3.6 | 3.7 | 3.2 | 3.7 | 3.5 | 3.7 | 3.4 | 3.1 | 2.9 | 2.7 | 3.0 | 2.7 | 2.8 | 2.8 | 2.7 |
| 81 | 3.7 | 3.6 | 3.5 | 3.7 | 3.6 | 3.3 | 3.3 | 3.4 | 3.3 | 3.0 | 3.0 | 2.8 | 2.8 | 2.8 | 2.7 | 2.6 | 2.8 |
| 84 | 3.6 | 3.6 | 3.7 | 3.8 | 3.8 | 3.5 | 3.1 | 3.4 | 3.3 | 3.2 | 2.9 | 2.9 | 2.9 | 2.8 | 3.0 | 2.9 | 2.6 |
| 87 | 3.3 | 3.2 | 3.2 | 3.3 | 3.3 | 3.0 | 3.1 | 3.0 | 3.0 | 2.6 | 2.5 | 2.7 | 2.5 | 2.6 | 2.6 | 2.7 | 2.4 |
| 90 | 1.8 | 0.8 | 1.1 | 1.5 | 1.2 | 1.1 | 0.7 | 1.5 | 1.2 | 1.3 | 0.7 | 0.0 | 0.9 | 1.2 | 0.9 | 0.5 | 1.0 |

Table 2b. Luminous intensity values, azimuth 190-350°

Signature:

_________________________________

Print Name:

GH JOHN

_________________________________

Date:

01-02-2013

_________________________________

Partner / Director
Duly authorised to sign on behalf of:

Photometric and Optical Testing Services LLP